## Slide 1: Small Sample Volumes are Small Problems
Kris Franklin, Chief Scientific Officer
Andréa Forsyth, PhD. Senior Scientist
Timothy J Dodson, Mol. Biologist & Chemist
Ashley Navarro, Medical Technologist

## Slide 2: Small Volume Testing- the Good, the Bad & the Ugly
Sample handling is unique to the matrix.
Matrices require special sample prep.
  No “one size fit all” option.
  “Dilute-and-shoot” won’t work.
Sensitivity of the instrument.
Inability to run sample multiple times
Depending on the matrix, may be less invasive than traditional testing.
Collection Time is shorter
Lower shipping costs
Smaller storage requirements
THE GOOD
THE BAD & UGLY
A recent literature search of American Chemical Society reveals 1642 publications within the last year of “small volume samples” with clinical aspects

## Slide 3: Endocrine Panel Objectives
Develop a less invasive way to get a baseline measurement on patient.
Allows physician to have a window into how a patient is responding to therapy, or determine there is a need of therapy or supplementation.
The ability to routinely monitor a patient to observe circadian changes or patterns in hormones
Remove the need for venous puncture allowing for point of care type collection device

## Slide 4: Endocrine Panel Challenges
Endogenous problems:
  Levels vary by time (day of the month, time of collection, etc.)
  Some hormones may be bound to proteins (SHBG, albumin, etc).
Top image: Endocrine Surgeon.co.uk. http://www.endocrinesurgeon.co.uk/index.php/adrenals-the-cortex-cortisol
(Accessed Feb 12, 2016).
Bottom Image: S H Ho Urology and Laparoscopy Centre. http://www.urologycentre.com.sg/hypogonadism_testosteronelevel.html (Accessed Feb 12 2016).

## Slide 5: Endocrine Panel Challenges - Specificity
Multiple Isobaric Pairs:
  Aldosterone, Cortisone
  11-deoxycortisol, 21-deoxycortisol, corticosterone
  DHEA, Testosterone, Estrone
  17a-hydroxyl-progesterone, 11-deoxycorticosterone
Hormone | Mass | 
11-deoxycorticosterone | 330.461 | 
17-alpha-hydroxyprogesterone | 330.461 | 
21-deoxycortisol | 346.46 | 
11-deoxycortisol | 346.461 | 
corticosterone | 346.47 | 
cortisone | 360.44 | 
cortisol | 362.46 | 
Hormone | Mass | 
estrone | 270.366 | 
estradiol | 272.382 | 
androstenedione | 286.409 | 
estriol | 288.38 | 
testosterone | 288.42 | 
dhea | 288.424 | 
progesterone | 314.46 | 

## Slide 6: Endocrine Panel Challenges
Sensitivity
Hormones are poor ionizers (no proton donor / acceptor group)
Clinically many requires greater sensitivity and several have a wide dynamic range
Cholesterol similar structure in a significantly higher concentration
Matrix Effects and Ion Suppression may have more significant effect on sensitivity than drug analytes.
American Society of Hematology. http://www.hematology.org/Patients/Basics/ (Accessed Feb 1, 2016).
Medical Assessment. http://medicalassessmentonline.com/terms.php?R=568&L=B (Accessed Feb 12, 2016).

## Slide 7: Sanis vs. Traditional Methods
350 μL sample is used to quantify 14 different hormones.
Capillary puncture is safe, simple, and gives reliable results.
No need for phlebotomist, nurse, etc.
Images from: Austin Community College. http://www.austincc.edu/mlt/phb/phb_equipment.html
(Accessed Nov 11, 2015).
Traditional methods use venous puncture and focus on only a few hormones in per prep or injection
Larger volume & additional Medical Professional required.

## Slide 8: LC-MS/MS & Robotic Sample Prep Methodology
Single injection
Semi-Automated Sample preparation
8 min run time with 8.5 min loop time
Positive /Negative ionization single run
R2 value of 0.99 or greater
CV less than 20%

## Slide 9: Sanis vs Other Laboratories
 |  | Sanis Biomedical |  |  | Lab Y |  |  | Lab Z | 
 | Min vol (mL) | Partnered w/other tests? | Turn around? | Min vol mL | Partnered w/other tests? | Turn around? | Min vol | Partnered w/other tests? | Turn around?
Cortisol | 0.35 | All analytes listed. | 2-3 d | 0.3 | yes (cortisone) | n. l. | 0.2 | yes (0.5 ml x 4) | n. l.
Cortisone | 0.35 |  | 2-3 d | 0.2 | yes (cortisol) | n. l. | 0.3 | yes (0.25 ml) | n. l.
Corticosterone | 0.35 |  | 2-3 d | 0.2 | yes (1.5 ml) | n. l. | 0.25 | yes (0.25 ml) | n. l.
11-Deoxycortisol | 0.35 |  | 2-3 d | 0.2 | yes | n. l. | 0.25 | yes (0.25 ml) | n. l.
21-Deoxycortisol | 0.35 |  | 2-3 d | n/a | n/a | n. l. | n/a | n/a | n. l.
11-Deoxycorticosterone | 0.35 |  | 2-3 d | 0.5 | no | n. l. | n/a | n/a | n. l.
17-alpha-Hydroxyprogesterone | 0.35 |  | 2-3 d | 0.5 | no | n. l. | 0.25 | yes (0.25 ml) | n. l.
Progesterone | 0.35 |  | 2-3 d | 0.3 | yes (2 ml) | n. l. | 0.25 | yes (0.25 ml) | n. l.
Estrone (E1) | 0.35 |  | 2-3 d | 1 | yes (2.2 ml) | n. l. | 0.25 | yes (0.8 ml) | n. l.
Estradiol (E2) | 0.35 |  | 2-3 d | 0.6 | yes (2.2 ml) | n. l. | 0.45 | yes (0.8 ml) | n. l.
Estriol (E3) | 0.35 |  | 2-3 d | 0.3 | yes (2.2 ml) | n. l. | 0.3 | yes (0.8 ml) | n. l.
Androstenedione | 0.35 |  | 2-3 d | 0.5 | yes (0.5 ml) | n. l. | 0.25 | yes (0.6 ml) | n. l.
Dehydroepiandrosterone | 0.35 |  | 2-3 d | 0.5 | yes (0.5 ml) | n. l. | 0.3 | yes (0.25 ml, or 1.1 ml) | n. l.
Testosterone | 0.35 |  | 2-3 d | 0.4 | yes (1 ml) | n. l. | 0.2 | yes (0.25 ml, or 0.6 ml) | n. l.

## Slide 10: Scientific Outline & Thought Process
  Can the method be harmonized to regulatory standards?
  Correlation study with CDC (testosterone)
  Would a doctor get the same result as a traditional method?
  Correlation study with Mt. Sinai (cortisol, testosterone)
  Can capillary prick replace the venous puncture?
  Correlation study with venous puncture versus capillary prick data.
  Can we get the same result reproducibly?
  Precision using external control (UTAK steroids mixtures; levels 1 and 3).
Does the method meet all the requirements of a clinical validation?
  Validation Guidelines based on CLSI guidelines

## Slide 11: Can the method meet CDC standards?
Chart: CDC (ng/mL)

Identity | 0 | 9
Allowable difference ±25% | 0 | 6.75 | 0 | 11.25
40 samples from the CDC tested Concentrations ranged from ~0.07 ng/ml to ~8 ng/ml
Chart: CDC (ng/mL)

Identity | 0 | 0
Mean (1.31%) | 0.013076 | 0.013076
95% LoA (-33.79% to 36.41%) | -0.33792 | -0.33792 | 0.36407 | 0.36407
Notes: 40 samples from the CDC tested multiple times.
  Concentrations ranged from ~0.07 ng/ml to ~8 ng/ml

## Slide 12: Would the doctor get the same result, compared to a traditional method?
Sought to compare data from two commonly measured hormones:
  Cortisol (n = 17)
  Testosterone (n = 22)
Samples provided by Mt. Sinai hospital (Toronto)

## Slide 13: Cortisol - Correlation Study with Mt. Sinai
Chart: Sanis (ng/mL)

Identity | 25 | 275
 | 46.805 | 153.988 | 134.564 | 108.945 | 185.235 | 54.397 | 66.666 | 264.223 | 198.064 | 36.869 | 233.972 | 68.057 | 153.479 | 251.891
Chart: 

Mean (8.2525)
95% LoA (-3.8937 to 20.3987)
Chart: (Target + Conc.) / 2

Identity | 0 | 0
 | 1.495 | 10.818 | 11.334 | 5.645 | 17.425 | 6.187 | 0.696 | 17.033 | 6.324 | 5.699 | -0.898 | 3.907 | 13.209 | 16.661
Mean (8.2525) | 8.2525 | 8.2525
95% LoA (-3.8937 to 20.3987) | -3.8937 | -3.8937 | 20.3987 | 20.3987
Chart: 
Frequency | 0 | 1 | 7 | 6 | 0

Labels

## Slide 14: Testosterone – Correlation Study with Mount Sinai
Chart: Mt. Sinai

Identity | 0 | 7
 | 2.378 | 2.997 | 0.11 | 5.515 | 0.097 | 0.09 | 0.09 | 4.906 | 6.297 | 2.897 | 3.78 | 3.347 | 0.981 | 1.141 | 0.996 | 6.383 | 5.927 | 1.732 | 1.699 | 1.936 | 0.436 | 0.469 | 0.124 | 4.724 | 2.031 | 0.174 | 4.188 | 3.868
Allowable difference ±25% | 0 | 5.25 | 0 | 8.75
Chart: 

Allowable difference ±25%
Chart: Target (ng)

Identity | 0 | 0
 | -0.20874 | -0.22915 | -0.15247 | -0.15204 | 0.20113 | 0.11445 | 0.11445 | -0.20291 | 0.02309 | -0.16017 | 0.09581 | -0.029715 | -0.18041 | -0.046738 | -0.16788 | -0.022054 | -0.091918 | -0.12077 | -0.13752 | -0.017213 | 0.3625 | 0.1725 | -0.31111 | -0.15643 | -0.18105 | -0.033333 | -0.073451 | -0.20738
Chart: 
Frequency | 0 | 5 | 16 | 5 | 2

Labels

## Slide 15: Can capillary prick compare to venous puncture?
Small scale study, looked at the entire panel.
Good correlation between venous and capillary measurements.
Critical to being able to replace current trends
Larger scale study planned.

## Slide 16: Cortisol (Vein vs. Capillary)
Chart: (Vein + Cap) / 2

Identity | 1 | 1
 | 1.1573 | 1.1246 | 1.0868 | 1.0758 | 0.90281 | 0.91507 | 0.83877 | 0.84705 | 0.9938 | 1.0198 | 1.018 | 0.99886 | 1.2797 | 1.1732 | 1.2136 | 1.2052 | 1.0564 | 1.0724 | 1.3128 | 1.3292 | 1.0837 | 1.0563 | 1.0193 | 1.0286 | 1.0456 | 1.0439 | 0.94945 | 0.93269 | 0.88088 | 0.89673 | 0.88617 | 0.86934 | 1.0178 | 1.093 | 0.94428 | 0.97964 | 0.87679 | 0.86523 | 1.1186 | 1.119 | 0.95809 | 0.90827
N | 42 | 
 | Minimum | Maximum
Vein | 4.2170 | 180.6270
Cap | 4.4260 | 179.5080
Correlation - r | 0.973 | 
Chart: Vein

 | 114.2 | 106.808 | 105.664 | 107.471 | 35.188 | 34.703 | 33.611 | 32.382 | 179.508 | 178.724 | 179.034 | 176.851 | 115.389 | 108.154 | 108.115 | 107.821 | 81.552 | 82.092 | 91.564 | 93.63 | 58.575 | 58.948 | 54.704 | 55.777 | 67.455 | 66.186 | 59.708 | 60.28 | 114.152 | 123.297 | 116.569 | 116.554 | 118.126 | 121.054 | 116.387 | 117.468 | 109.162 | 108.133 | 4.717 | 4.862 | 4.481 | 4.426
Passing-Bablok fit (y = 1.512 + 0.9959 x) | 5.7115 | 181.405
Allowable difference ±20 | -20 | 180 | 20 | 220

## Slide 17: Testosterone (venous vs. capillary data)
Chart: Vein

 | 3.72 | 3.717 | 3.799 | 3.819 | 0.275 | 0.269 | 0.257 | 0.242 | 0.261 | 0.284 | 0.316 | 0.324 | 3.839 | 3.691 | 3.82 | 3.88 | 2.925 | 2.881 | 2.874 | 2.913 | 0.263 | 0.322 | 0.275 | 0.278 | 4.172 | 4.178 | 4.16 | 4.147 | 4.455 | 4.3 | 3.992 | 3.978 | 0.16 | 0.138 | 0.217 | 0.237 | 0.292 | 0.28 | 4.285 | 4.385 | 4.45 | 4.377
Passing-Bablok fit (y = 0.05273 + 0.9254 x) | 0.10548 | 5.2053
Allowable difference ±0.75 | -0.75 | 5.25 | 0.75 | 6.75
N | 42 | 
 | Minimum | Maximum
Vein | 0.0570 | 5.5680
Cap | 0.1380 | 4.4550
Correlation - r | 0.987 | 
Passing-Bablok fit |  |  |  |  |  | 
Equation | Cap = 0.05273 + 0.9254 Vein |  |  |  |  | 
Parameter | Estimate | Bootstrap 99% CI |  |  |  | 
Intercept | 0.05273 | 0.004135 | to 0.1402 |  |  | 
Slope | 0.9254 | 0.8884 | to 0.9596 |  |  | 
CI based on 99 bootstrap samples. |  |  |  |  |  | 

## Slide 18: Can the Results be consistently Reproduced?
Externally Prepared UTAK QC material at clinical levels
Measured over the course of a month to determine stability of QC material
Confidence that Results will be the same over time

## Slide 19: Cortisol Precision with External UTAK Serum
LOW QC
HIGH QC
Chart: Observation #

 | 22.116 | 23.472 | 22.562 | 23.068 | 23.683 | 23.538 | 23.279 | 23.304 | 24.173 | 22.933 | 23.915 | 24.685 | 23.859 | 24.479 | 24.125 | 24.211 | 25.704 | 25.386 | 25.292 | 22.886 | 26.903 | 27.249 | 26.054 | 26.441 | 27.965 | 29.655 | 29.986 | 29.643 | 30.527 | 29.59 | 30.449 | 30.211 | 30.803 | 29.933
Mean | 25.9435 | 25.9435
Chart: Observation #

 | 103.328 | 104.263 | 111.756 | 107.708 | 120.722 | 112.299 | 110.292 | 139.245 | 111.907 | 116.016 | 110.059 | 108.456 | 110.298 | 111.091 | 108.977 | 113.038 | 113.485 | 116.516 | 117.127 | 107.622 | 134.591 | 125.302 | 130.275 | 140.628 | 153.565 | 157.205 | 153.79 | 151.444 | 160.105 | 149.159 | 150.745 | 144.408 | 146.362 | 156.167
Mean | 126.7044 | 126.7044
LOW QC |  |  | 
 | Avg. | Std. Dev. | CV
Conc. | 25.944 | 2.8148 | 10.850
Ratio | 4.027 | 0.1942 | 4.822
ISTD | 1,811,550 | 212643.3 | 11.738
HIGH QC |  |  | 
 | Avg. | Std. Dev. | CV
Conc. | 126.704 | 18.8101 | 14.846
Ratio | 19.578 | 1.1336 | 5.790
ISTD | 1,503,191 | 152952.1 | 10.175

## Slide 20: Testosterone Precision & Accuracy UTAK Serum 2 Levels
Chart: Observation #

 | 4.36 | 4.4 | 4.68 | 4.616 | 4.84 | 4.635 | 4.484 | 5.763 | 4.627 | 4.573 | 4.679 | 4.759 | 4.753 | 4.517 | 4.529 | 4.297 | 4.589 | 4.435 | 4.682 | 4.68 | 4.887 | 4.506 | 5.101 | 5.135 | 4.931 | 5.145 | 5.192 | 5.009 | 5.408 | 4.787 | 4.688 | 5.218 | 5.058 | 5.265
Mean | 4.8008 | 4.8008
Chart: Observation #

 | 0.944 | 0.936 | 0.945 | 0.965 | 0.946 | 0.954 | 0.898 | 0.984 | 0.927 | 0.978 | 0.972 | 1.02 | 0.977 | 0.966 | 0.986 | 0.984 | 0.99 | 0.983 | 0.964 | 0.939 | 0.987 | 0.948 | 0.977 | 0.94 | 1.073 | 0.959 | 0.961 | 1.031 | 0.978 | 0.987 | 1.028 | 1.027 | 1.016 | 1.003
Mean | 0.97568 | 0.97568
LOW QC |  |  | 
 | Avg. | Std. Dev. | CV
Conc. | 0.976 | 0.0345 | 3.538
Ratio | 0.131 | 0.0068 | 5.188
ISTD | 10,171,862 | 2888714 | 28.399
HIGH QC |  |  | 
 | Avg. | Std. Dev. | CV
Conc. | 4.801 | 0.3272 | 6.816
Ratio | 0.645 | 0.0534 | 8.278
ISTD | 11,491,413 | 3326020 | 28.944

## Slide 21: Fully Designed Clinical Validation for Lab Developed Test (LDT)
Linearity
Sensitivity
Carryover
Ion Suppression
Limit of Quantification
Reproducibility

## Slide 22: Cortisol Linearity
Chart: X Variable 1 Residual Plot
| 0.39500000000000002 | 0.378 | 0.38300000000000001 | 0.41499999999999998 | 0.38500000000000001 | 0.41899999999999998 | 0.45900000000000002 | 0.34899999999999998 | 0.56000000000000005 | 0.54800000000000004 | 0.65 | 0.61499999999999999 | 0.73099999999999998 | 0.749 | 0.63800000000000001 | 0.65400000000000003 | 0.77200000000000002 | 0.72 | 0.72399999999999998 | 0.63100000000000001 | 1.3520000000000001 | 1.298 | 1.2350000000000001 | 1.26 | 1.3260000000000001 | 1.339 | 1.323 | 1.2430000000000001 | 1.38 | 1.3859999999999999 | 2.5910000000000002 | 2.3849999999999998 | 2.9329999999999998 | 2.4350000000000001 | 2.8439999999999999 | 2.431 | 2.4630000000000001 | 2.9460000000000002 | 2.5760000000000001 | 2.4790000000000001 | 5.3220000000000001 | 4.9530000000000003 | 4.8010000000000002 | 5.17 | 5.0140000000000002 | 5.0940000000000003 | 5.3109999999999999 | 5.2279999999999998 | 5.31 | 5.1120000000000001 | 8.5440000000000005 | 9.8789999999999996 | 10.151999999999999 | 9.2490000000000006 | 8.8620000000000001 | 9.4529999999999994 | 10.311 | 9.8870000000000005 | 9.4969999999999999 | 9.33
 | -1.8107e-05 | -6.089e-05 | 3.6423e-05 | 8.1643e-05 | -4.0832e-05 | -2.6567e-05 | 6.6134e-05 | 5.0572e-06 | 7.8309e-05 | 8.9384e-06 | 9.9352e-06 | -6.7853e-05 | -5.7841e-05 | 3.7646e-06 | -4.8135e-05 | -4.6375e-05 | -3.5518e-05 | 9.6512e-05 | -2.9798e-05 | 9.6907e-05 | -3.4457e-05 | 2.2727e-05 | 3.0109e-05 | -1.9329e-05 | 2.5577e-06 | 3.105e-05 | -5.071e-05 | 4.2489e-05 | 6.3737e-06 | -6.1091e-05 | 2.3043e-05 | -1.9995e-05 | -8.4552e-06 | 8.1131e-05 | 5.9941e-05 | 3.5441e-05 | -2.9039e-05 | -6.5962e-05 | -3.2294e-05 | -8.2786e-06 | -1.0587e-06 | -1.0968e-05 | 8.092e-07 | -2.8282e-05 | -1.9194e-05 | -6.5393e-05 | 4.7294e-05 | -6.0775e-05 | -6.6129e-05 | 6.9212e-05 | -1.6704e-05 | 7.7358e-05 | 4.7053e-06 | 2.1171e-05 | 5.9656e-05 | -2.6635e-05 | -6.2114e-05 | -7.7262e-05 | 5.7956e-05 | 3.9395e-05

## Slide 23: Testosterone Linearity
Chart: X Variable 1 Residual Plot
| 0.2 | 0.189 | 0.218 | 0.21299999999999999 | 0.191 | 0.21099999999999999 | 0.19 | 0.19900000000000001 | 0.193 | 0.19800000000000001 | 0.33100000000000002 | 0.374 | 0.35399999999999998 | 0.374 | 0.32100000000000001 | 0.33900000000000002 | 0.35199999999999998 | 0.36099999999999999 | 0.373 | 0.36099999999999999 | 0.70799999999999996 | 0.64700000000000002 | 0.623 | 0.65800000000000003 | 0.70699999999999996 | 0.68200000000000005 | 0.63600000000000001 | 0.61499999999999999 | 0.69 | 0.68400000000000005 | 1.175 | 1.1080000000000001 | 1.5509999999999999 | 1.2090000000000001 | 1.569 | 1.2889999999999999 | 1.35 | 1.534 | 1.2929999999999999 | 1.3320000000000001 | 3.008 | 2.8220000000000001 | 2.2450000000000001 | 2.331 | 2.58 | 2.8740000000000001 | 2.8079999999999998 | 2.8650000000000002 | 2.4870000000000001 | 2.516 | 3.84 | 5.165 | 5.4630000000000001 | 4.6950000000000003 | 4.5449999999999999 | 4.5270000000000001 | 4.4409999999999998 | 4.484 | 5.0330000000000004 | 4.1070000000000002
 | 4.3725e-05 | 5.0901e-05 | -5.848e-05 | -1.7273e-05 | 2.0782e-06 | 3.625e-05 | -4.686e-05 | -4.1713e-05 | -4.1145e-05 | -9.3517e-06 | -1.7321e-05 | -7.0616e-06 | 4.3867e-05 | -3.0762e-05 | 2.8294e-05 | 3.4188e-05 | -6.121e-05 | -3.3163e-05 | -3.96e-05 | 3.7837e-05 | -1.8276e-05 | -4.2029e-05 | -6.3255e-05 | 4.5895e-05 | 1.8985e-05 | 4.9921e-05 | 5.9146e-05 | 5.1636e-05 | 6.463e-05 | -4.1502e-05 | 2.3041e-05 | -1.3044e-05 | 1.1547e-05 | -6.2647e-05 | -4.7586e-06 | 1.5439e-05 | -1.9708e-05 | 2.8391e-05 | -5.2407e-05 | 4.1598e-05 | 6.7258e-05 | -1.4918e-05 | -4.9679e-05 | -2.7361e-05 | 3.7746e-05 | 4.9088e-05 | -4.2458e-05 | -1.1759e-05 | 6.7658e-05 | 6.1277e-05 | -6.7644e-05 | 1.3975e-05 | -4.9528e-05 | -1.2158e-05 | 5.6055e-05 | -6.3912e-07 | -5.7957e-05 | -1.7798e-05 | 5.5882e-05 | -2.2843e-05
Specificity | AVG | St Dev | CV
Ret Time | 3.386 | 0.008 | 0.245
MRM Ratio | 89.272 | 3.086 | 3.457

## Slide 24: Applications of Small Volume Blood
Direct Effect of Therapies effecting Hormone imbalance
  Dexamethasone Suppression
Routine Monitoring for Environment Stress Factors
  Cortisol & Testosterone
  Effects of Diet & Supplementation
Notes: 2 application of point of care small volume blood draw

## Slide 25: Dexamethasone Suppression Effect
Two patients enrolled in our test had been given dexamethasone shots.
 | Conc. (ng/ml) | Average (ng/ml) | Clinical Range (ng/ml)
Patient 1 |  |  | 
Cortisol | 9.853 | 130.00 | 70 - 280 (a.m.)
Cortisone | 2.003 | 26.40 | 12 - 35 (a.m.)
Patient 2 |  |  | 
Cortisol | 4.21 | n/a | 70 - 280 (a.m.)
Cortisone | 0.35 | n/a | 12 - 35 (a.m.)

## Slide 26
Can quantify the effects of long hours of work or over exhaustion.
Marker for overtraining can be established.
Food and supplements that can lower cortisol can be tested.
Can give a view for whether or not an athlete’s diet or supplements are affecting their health.
For supplements:
  Are they giving me the boost that I want/need?
  Are they effecting me in other ways?
Effects of Stress and Supplementation

## Slide 27: Summary
Small volumes present unique, but not insurmountable challenges.
  Good sample prep is key; regardless of matrix.
Small sample volumes can be used to search for multiple analytes with accuracy and sensitivity.
Ultimate goal is always to provide accurate information in a minimally invasive way to the patient and primary care provider.

## Slide 28: The Brighter Path through Innovative Testing –Sanis Biomedical
Small sample serum toxicology panel
  Allows for Therapeutic Drug Monitoring
  Allows application of Goodwin & Gillman ranges in the clinical space
Endogenous markers and therapeutic drug monitoring
  Utilziing Small volume Chemistry to obtain more health related information quicker.

## Slide 29: Toxicology Panel 60 - 80 Analytes
6 General Classes:
  Stimulants
  Barbiturates
  Muscle Relaxants
  Opioids
    Synthetic
    Natural
  Psychiatric Medications
    SSRI
    TCA
    Benzodiazepines
    Anti-Psychotics
  Substances of Abuse
    Bath Salts
    Synthetic Cannabinoids
    Cannabinoids

## Slide 30: Endogenous Biomarkers and Therapeutic Drug Monitoring
5 Main Classes:
  Inflammatory markers
  Neurotransmitters
  Therapeutic Drug Monitoring
  Thyroid Markers
  Vitamins and Amino Acid Metabolism
Notes: Inflammatory markers
Thyroid markers
Vitamins & Amino acid metabolism
Neurotransmitters
TDM

## Slide 31: Endogenous Biomarkers and Therapeutic Drug Monitoring
Entry | Name | Class
1 | Uric acid | IM
2 | Lactic acid | IM
3 | Homocysteine | IM
4 | Bilirubin | IM
5 | Vanillylmandelic acid | NT
6 | Homovanillic acid | NT
7 | Mandelic acid | NT
8 | Valproic acid | TDM
9 | Gabapentin | TDM
10 | Salicylates | TDM
11 | Phenobarbital | TDM
12 | Naproxen | TDM
13 | Ibuprofen | TDM
Entry | Name | Class
1 | T-3 | THY
2 | Reverse T-3 | THY
3 | T-4 | THY
4 | Tyrosine | V+AA
5 | 3,4-dihydroxy-L-Phenylalanine | V+AA
6 | L-Tryptophan | V+AA
7 | L-Histidine | V+AA
8 | L-Phenylalanine | V+AA
9 | Niacin | V+AA
10 | Folic acid | V+AA
11 | Biotin | V+AA
12 | Pyridoxine hydrochloride | V+AA
13 | L-Ascorbic acid | V+AA
Notes: Inflammatory markers
Thyroid markers
Vitamins & Amino acid metabolism
Neurotransmitters
TDM

## Slide 32: External Acknowledgment
Dr. Eleftherios P. Diamandis, MD, PhD, FRCP(C),FRSC
  Mt. Sinai Clinical Hospital
Centers for Disease Control
Trinity (Vein vs. Capillary)
Biotage
Shimadzu

## Slide 33: End
Questions?